## Slide 1: Soft X-ray hybrid detectors using LGAD sensors
A. Bergamaschi, M. Anastasi, M. Brückner, M. Carulla, R. Dinapoli, S. Ebner, S. Emiliani, K. Ferjaoui, E. Fröjdh, V. Gautam, D. Greiffenberg, S. Hasanaj, V. Hinger, M. Hürst, V. Kedych, T. King, S. Li, C. Lopez-Cuenca, L. Lunin, A. Mazzoleni, D. Mezza, S. Moses, K. Moustakas, A. Mozzanica, J. Mulvey, M. Müller, C. Ruder, B. Schmitt, S. Silletta, D. Thattil, X. Xie, J. Zhang
NICE 2026, Latisana (Italy)
Thursday 21st May 2026
Center for Photon Science, Paul Scherrer Institute (Switzerland)

## Slide 2: Why detector development?
Sample
Data
Light Source
Optics
Detector
Synchrotrons, XFELs
Bending magnets
Insertion devices
Beamline
Monochromator
Mirrors
Preparation
Environment
Methods
Position sensitive
Energy dispersive
SCIENCE

## Slide 3: Detecting soft X-rays is hard!
Shallow absorption of low energy X-rays
Small signal comparable to electronic noise
Thickness of Si unsensitive
entrance window (um)
QE (%)
800 eV single photon spectrum with JUNGFRAU
Hinger et al.,
https://doi.org/10.1088/1748-0221/17/09/C09027

## Slide 4: State-of-the-art Soft X-ray Detectors Monolithic detector
Small pixels
Low noise
Thin sensitive region
Limited dynamic range
Large periphery
Radiation damage
CCDs
Slow readout
Shutter required
Cooling
sCMOS
Moderate frame rates
Rolling shutter
Evolving technology
Dedicated developments: pnCCD, vfCCD, PERCIVAL…

## Slide 5: Hybrid detectors
Semiconductor sensor
  Direct conversion
  Thickness, material and technology can be optimized for X-ray energy
Versatile CMOS readout chip
  Fast readout (kHz)
  Radiation hardness (Mrad)
Bump bonding
  Minimum pixel pitch ~ 20 μm
  Noise due to input capacitance
Large area by tiling

## Slide 6: Hybrid detector development at PSI

## Slide 7: Hybrid detectors for soft X-rays
Technology maturity
Suitability for soft X-ray applications
(synchrotrons/FELs)
BSI
sCMOS
CCD
MCP
MAPS FD-MAPS
Hybrid LGAD
Courtesy A. Dawiec, SOLEIL
Courtesy J. Zhang, PSI

## Slide 8: Quantum efficiency
Nitrogen and Oxygen
K-edge
  Shallow absorption of soft X-rays
  Reduce the thickness of the absorbers on top
  Decrease the charge recombination on the surface
  Reduce the charge recombination in the implants
n+
n-
Al
p+
Optimized
Al
x
x x
ASIC
n+
n-
Al
p+
Al
x x x x x x x x x x x x x x x x
ASIC
Standard
Hard X-ray sensors
Optimized implant
Passivated surface
Thin passivation
Desjardins, et al. JSR 2020.
Send, et al. NIMA 2013.
J. Zhang et al., JINST (2023).
Carulla, et al. Sensors. 2024; 24(3): 942.
SiO/N

## Slide 9: Quantum Efficiency Extended to the UV Energy Range
The QE is determined by the absorption of the passivation dielectric and the charge collection efficiency
Radiation hardness tested in direct beam
Second batch
92 eV
Beam test @ Soleil
Courtesy S. Li, PSI
Courtesy J. Zhang, PSI
Notes: High QE for 1. eUV photon 2.Rixs with shallow incidence angle 3. 10 keV electron (~nm)

## Slide 10: Low Gain Avalanche Diodes
Use internal amplification to improve SNR
  Like APDs, but no dark counts and can be segmented
Soft X-ray
Gain layer
ASIC readout
Amplification
Charge multiplication induced by impact ionization

## Slide 11: Segmented LGAD for time resolved STXM
LGAD originally developed for time resolution
Segmentation possible to avoid pile-up
Promising results comparable to APD to separate electron 500MHz bunches at synchrotrons
Dedicated production for soft X-ray detection required for optimal performance
Courtesy S. Li
In collaboration with M. Weigand

## Slide 12: From high energy physics to photon science
LGAD with gain layer on readout side
iLGAD with gain layer on sensor backside
Pencil beam scan for LGAD (measurement)
Gain variation for iLGAD (simualtion)
multiplication
multiplication
multiplication
Courtesy J. Zhang
Electron collection
Standard in particle physics
Limited fill factor
Uniform gain
Hole collection
Gain dependent from the depth of absorption
Complex fabrication

## Slide 13: Depth-of-absorption dependence
900 eV spectrum measured with MÖNCH
Mh
n+
gain layer
X-rays
depth
Me
e-
h+
Detector output [eV] ∝ gain
h+ multiplication
peak
e- multiplication
peak
Double peak due to the different impact ionization factor of e- and h+
Courtesy J. Zhang

## Slide 14: Gain Layer Optimization
Double peak features due to e/h-initiated mulitplication
Shallow gain design benefits from larger probability of e-initiated multiplication
Shallow and thin gain design benefits from higher probability of e--initiated multiplication
  Me/Mh ratio depends on the electric field
  Currently lower Me for shallow gain, due to missing impact ionization model for simulations
    Corrected in the current batch
Mh
n+
gain layer
X-rays
attenuation
depth
Mh
n+
gain layer
X-rays
depth
900 eV
Standard gain design
Shallow gain design
Me
Me
900 eV
Standard gain design
Shallow gain design
Different gain layer designs
Courtesy J. Zhang

## Slide 15: Single Photon Detection below 1 keV
Double peak features due to e/h-initiated mulitplication
Shallow gain design benefits from larger probability of e-initiated multiplication
390 eV
h+ initiated multiplication peak
e- initiated multiplication peak
SNR > 5 for photon energy > 400 eV (both shallow and standard gain design)
  Similar results for MÖNCH and JUNGFRAU
Standard sensor
SNR
SNR of e- multi. peak
SNR of h+ multi. peak
X-ray Energy (eV)
Weighted SNR
Higher gain and shallower profile to achieve lower energies
Courtesy J. Zhang

## Slide 16: Single Photon Counting using iLGADs
4x4 cm2 sensor (512x512 pixels)
  Largest iLGAD sensor to date!
  Low yield!
Single photon counting readout chip
Cooling (-20 to -40 deg) improves performance
  Lower leakage current
  Higher gain
Eiger iLGAD quad mounted in the SOPHIE chamber
Courtesy E. Frojdh

## Slide 17: Single photon counting down to 500 eV
M = iLGAD sensor gain
Baruffaldi, F., et al. Single-photon counting pixel detector for soft X-rays. Commun Phys 8, 321 (2025). https://doi.org/10.1038/s42005-025-02240-9
Notes: Explane DQE(0)

## Slide 18: Ptychography at the Fe L3-edge
LGAD Eiger (standard gain layer)
at 712.5 eV Spatial resolution ~8 nm
Amplitude
Phase
Baruffaldi, et al.
Single photon counting pixel detector for soft X-rays
Communications Physics 8, 321 (2025).
Noiseless
High dynamic range
Fast frame rate
Courtesy T. Butcher

## Slide 19: Counting detectors
“Noiseless”
  Minimum detectable energy
  Fluorescence suppression
Stable, rad-hard, user friendly
Count rate limited
Photons
Analog output
19

## Slide 20: No counting at pulsed sources
1 pulse
=
1 count
Log photons
Analog output

## Slide 21: … or charge integrating detectors?
Photons
Analog output
Single photon counting
Detector output
Charge integrating
Photons are like chocolates…

## Slide 22: Dynamic range of charge integrating detectors
1-10000
photons/bunch
0.01-100
kg/box
High gain
High precision
Medium gain
Medium precision
Low gain
Low precision
Adaptive gain
Adaptive precision

## Slide 23: PSI Detector portfolio
 | Microstrip 1D | Pixel 2D | 
 | 25/50 µm pitch | 75 µm pitch | 25 µm pitch
Single photon counting |  |  | 
Charge integrating |  |  | 
MYTHEN
GOTTHARD
MÖNCH
JUNGFRAU
EIGER

## Slide 24: Magnetic SAXS at laser plasma source
15'000 frames (150 s)
Courtesy L. Lunin
MÖNCH 400x400 25µm pitch pixels (1x1 cm2)
High SNR improves single photon discrimination
Single shot imaging possible
in collaboration with
D. Schick et al.

## Slide 25: Interpolation exploiting charge sharing
With fast frame rates and low fluxes
  Interpolation exploiting charge diffusion
  Super-resolution ~µm
Native resolution (25 µm pitch)
Interpolated to 1 µm bins
Cartier et al. (2016). J. Synchrotron Rad. 23, 1462-1473.
100 µm
LGAD sensor @ 500eV
25 µm pixels
2.5 µm bins
unpublished
Standard sensor @ 10 keV

## Slide 26: JUNGFRAU with rectangular pixels for RIXS
JUNGFRAU + iLGAD enables high-speed, high-resolution soft X-ray detection
QE > 80% @ 250 eV with Thin Entrance Window
SNR > 5 for single photons down to 400 eV with iLGADs
Rectangular pixels for interpolation
4 cm
4 cm
25 µm
75 µm
75 µm
Implant
Metal
X-ray photon
Interpolation direction
225 µm
Courtesy V. Hinger

## Slide 27: First RIXS test at SwissFEL
Good QE and SNR down to 850 eV
Spatial resolution after interpolation ≈ 3 µm
Courtesy V. Hinger,
S. Silletta
From focus beam scan at POLLUX
SLS
In collaboration with E. Paris et al.

## Slide 28: Pump-Probe RIXS at 50 kHz
Pump-probe experiment at 50 kHz at EuXFEL
  10 us between pulses thanks to16 JUNGFRAU storage cells
  Alternating pumped-unpumped pulses
Observed transient effect on sample
  Reduced emission at the end of train
  The effect increases with the beam intensity
Courtesy M. Ramilli, EuXFEL
28
In collaboration with L. Le Guyader et al.

## Slide 29: Challenges of iLGAD
Low yield (pixel and device)
Non-uniformities of the gain layer
Performance degradation at low energies
  Holes multiplication is predominant
Low quantum efficiency in grazing incidence
  Absorption in the passivation and hole multiplication merged with noise
29
Courtesy V. Hinger
Gain

## Slide 30: LEAPS-TECH
Several possible technologies to be explored
  Time, money
Need for a better understanding of features
  Improve simulations
  Characterization with X-rays
Collaborative effort required!
Courtesy G. Paternoster
Courtesy A. Bähr

## Slide 31: Outlook
LGADs could transfer to soft X-rays all the advantages of hybrid detector for hard X-rays
  First results already demonstrated
  The development of a reliable technology will take a few years
iLGADs allow single photon detection down to 400 eV
  New sensor batch under test: down to the carbon edge?
Alternative LGAD technologies might be more suitable for soft X-ray detection (or EUV)
  International effort through LEAPS, but also in the USA, UK, China…
Tests for different applications
  compare with advancing CMOS imagers and other technologies
  Input from users is more than welcome!

## Slide 32
Thanks for your attention
M. Anastasi, M. Brückner, M. Carulla, R. Dinapoli, S. Ebner, S. Emiliani, K. Ferjaoui, E. Fröjdh, V. Gautam, D. Greiffenberg, S. Hasanaj, V. Hinger, M. Hürst, V. Kedych, T. King, S. Li, C. Lopez-Cuenca, L. Lunin, A. Mazzoleni, D. Mezza, S. Moses, K. Moustakas, A. Mozzanica, J. Mulvey, M. Müller, C. Ruder, B. Schmitt, S. Silletta, D. Thattil, X. Xie, J. Zhang
S. Finizio, B. Watts, J. Raabe,
E. Paris, E. Razzoli, L. Pattey
A. Bisht, M. Boscardin,
M. Centis Vignali, G. Paternoster
M. Weigand
D. Schick, P. Wernicke, M. Matten
T. Butcher
N. Duarte, L. Le Guyader, M. Ramilli, M. Turcato, et al.

## Slide 33: Several soft X-ray detectors
Different soft X-ray sensors
Parameter | Andor iXon Ultra 888 (EMCCD) | Andor Neo 5.5 sCMOS | Princeton Instruments PIXIS 1024BR (CCD) | Hamamatsu ORCA-Fusion (CMOS) | pnCCD (PNSENSOR) | DSSC (EuXFEL) | PERCIVAL (DESY) | Hybrid X-ray detector (PSI, Jungfrau)
Quantum Efficiency (QE) @250 eV | 55% (400 nm UV) | 70% (400 nm UV) | < 42% (400 nm UV) | 65% (400 nm UV) | 80% | 52% | 90% | 80%
Read Noise (e-) | < 1 | 1.4 | 4-10 | 1.4 | 10.5 (hg) | 40-60 (miniSDD) 10 (DEPFET) | 16 | 3 - 5 (LGAD)
Frame Rate (FPS) | 26 | 30 (full frame) 100 (burst) | < 2 | < 100 | < 100 | 4.5 MHz (burst) | < 120 | 2 - 10 k
Dynamic Range (e-) | 80 k | 30 k | 100 k | 15 k | 1.6 M | 1.1 M | 3.5 M | 3.44 M (gain=10)
Pixel Size (μm) | 13 x 13 | 6.5 x 6.5 | 13 x 13 | 6.5 x 6.5 | 75 x 75 | 204 x 236 | 27 x 27 | 75 x 75 15 x 375
Pixels | 1024 x 1024 | 2560 x 2160 | 1024 x 1024 | 2304 x 2048 | 512 x 1024 | 128 x 256 | 1484 x 1408 | 512 x 1024
Sensor Area (mm²) | 13.3 x 13.3 | 16.6 x 14.0 | 13.3 x 13.3 | 14.9 x 13.3 | 38.4 x 76.8 (2 side buttable) | 30 x 62 (tiable) | 40 x 38 | 38.4 x 76.8 (tilable)
commercial
Facility / institute driven development
Courtesy J. Zhang, PSI

## Slide 34: The LGAD flavors
development @ PSI & FBK
 | LGAD | inverse-LGAD (iLGAD) | trench-isolated LGAD (TI-LGAD) | AC-coupled LGAD (AC-LGAD) | deep junction LGAD (DJ-LGAD) | ideal LGAD
cross section |  |  |  |  |  | ?
process* | standard | double-sided | stepper | standard | epi-growth | standard
complexity | low | low | medium | high | high | low
collected charge | e- | h+ | e- | e- (bipolar) | e- | e- or h-
readout | DC | DC | DC | AC | DC | DC
non-gain region (fill factor) | > 40 um low | 0 um 100% | 6-7 um high | 0 um 100% | 0 um 100% | 0 um 100%
interpolation for fine resolution | no | yes | no | yes | yes | yes
gain depends on absorption depth | no | yes | no | no | no | no
detection area | large | medium-large | small | medium-large | medium-large | large
multiplication of surface current | no | yes | no | yes | yes | no
risk/yield | low/good | medium/medium | medium/medium | medium/medium | high/low | low/good
Considered important for experiments
Courtesy J. Zhang, PSI

## Slide 35: Shallow gain vs. standard gain
Shallow gain design
standard gain design
e-
h+
e-
h+
e-
e-
h+
h+
Courtesy J. Zhang, PSI