## Slide 1: Photonic RF phase monitor
Russell Wilcox
Qiang Du, Tianhuan Luo, Shreeharshini Dharanesh Murthy, Will Wright, Alessandro Ratti
LLRF Workshop
October 14, 2026

## Slide 2: Sensor cable length drift affects precision LLRF control
LCLS-II LLRF control:
20m coax, 0.185 ppm /°C, time drift: 12 fs / °C
Precision feedback requires time-stabilized measurement
Measured out-of-loop RMS time jitter: 3.4 fs (0.1Hz – 5kHz)

## Slide 3: We propose time sensing via interferometry, and phase modulator transducer
Need stable (or measured) optical delay to sensor
  Popular: stabilized sampling pulses to EO transducer
    [M.K.Czwalinna, LLRF Workshop 2024]
    One sample per optical pulse
Proposed: CW to EO transducer, return modulated light
  Continuous data
  No bias, phase modulation is linear with voltage
  No electronics in radiation environment
Use optical interferometry for delay measurement
  Detect time delay with high precision:
    [R. Wilcox et al, Opt. Lett. 34, p3050 (2009)]
  Can separate carrier and modulation information
Use coherent fiber telecom components
  Optical phase detectors are mature tech:
    [K. Kikuchi, IEEE JLT 34, p157 (2016)]
CLK
Ampl. mod.
(MZM or SGI)
partial
reflector
time-stabilized link
RF
bias
electronics
pulses
Phase mod.
circulator
time-measured link
RF
CW
e.g. EuXFEL:
This talk:

## Slide 4: System block diagram
Phase modulator needs no bias or electronics
Interferometry: sub-fs Ddelay measurement
Coherent telecom devices: advanced functions, well-developed, low cost
  Doesn’t need LO sync or clock recovery
laser
DSP
phase mod
cavity
LLRF controls
ICR
freq. shift
LO
sig.
Integrated
Coherent
Receiver
added optical phase
adds to
pol. control
temperature
controlled
enclosure
time drift:
[Fujitsu FIM24725]

## Slide 5: Diagram of demodulation process
HPF
LPF
divide by
time shift
X, Y amplitude
phase
ICR
I
Q
ADC
subtract
group/phase correction
Drift corrected
Proportional to sensor voltage
RF amplitude is derived from polar phase, not amplitude
Works for any waveform, e.g. pulses
()+
()+
phase
time
phase
phase
time
time
rectangular to polar
known
RF phase error data,
used for feedback
multiply by
polarization calculation
X, Y IF phase
RF amplitude

## Slide 6: In principle, optical time delay correction uncertainty <1fs
No assumptions about what caused delay
(temp, air pressure, strain)
each wave is 5fs, 1˚is 14as
stability <3x10-10 or <150as (Rb locked laser)
the problem
accuracy <3x10-8 or 1as
[e.g. Radiantis “Locked Laser”]
Scheme: track changes in number of roundtrip waves, 2DN via IF phase
  Calculate phase delay Dtp and multiply by correction to get group delay Dtg
What intrinsic optical accuracy, given L = 100m, DT = 10˚C?
known to 3x10-8 (0.014as)
known to 10-4 (50as)
with 10-4 thermal correction (50as)
known to 10-4 (50as)
with 10-6 thermal correction (1as)

## Slide 7: Simulations show effects of errors
ANSYS Lumerical Interconnect model:
laser
frequency
shifter
phase
modulator
optical
hybrid
differential
photodiodes
I and Q
output
RF signal
input
Use simulation to estimate effect and
magnitude of error, compare with experiment
Ideal plot with photodetection noise
I
Q
Phase error in 90-degree optical hybrid
I
Q
I and Q offset 4% of an IF period, and with different optical phase delays
Q
I

## Slide 8: First experiment, no transport fiber
DSP
ICR
freq. shift
LO
sig.
Øpeak
1.12rad (observed)
≈ 2.5V
Power to modulator = 16dBm
detecting RF amplitude (optical phase subtracted):
I /Q, I (t), Q (t), showing optical phase drift
with RF = 110MHz, IF = 55MHz:
phase mod
laser
Xilinx ZCU208
7V (spec)
Amplitude measurement is insensitive to optical power variations

## Slide 9: Polarization modulation
IQ plots for one polarization:
experiment
simulation
constant
time-varying
EO phase modulator
X
Y
Output projected on X and Y axes of ICR
ICR sees:
mod depth
V,H dephasing
If polar amplitude is then that should fix it
PV
PH
PV
PH
40% p-p modulation even with PER=20dB, considered good.

## Slide 10: Polarization correction
ØXn - ØYn = 165˚ (phase difference)
AX/AY = EXsig/EYsig = 1.8 (amplitude ratio)
Complete polarization state is known
AX,ØX3
AX,ØX2
AX,ØX1
radius = 18k
X
AY,ØY1
AY,ØY2
AY,ØY3
radius = 10k
Y
DSP
phase mod
ICR
pol. control
Maintain polarization at modulator by compensating fiber depolarization without dithering
constant return state, fixed by polarizer
some constant state initially set when signal is maximized
linear
polarizer
polarization varies only within here
manual controller
(equalize X,Y)
USB connected 3-waveplate rotator:
sent state is constant
Three measurements of RF phase, not drift corrected:

## Slide 11: Short term data, 4GHz sample rate
1/IF
1/IF
slope
is IF
RF = 110MHz
IF = 55MHz
RF = 1GHz
IF = 55MHz
polarization
phase shift
time, µs
time, µs
IF phase, rad
polar phase:
I
Q
RF cycle
I vs Q (1GHz)
IF phase v time
LPF gives optical drift
two readings 0.5sec apart:
One datum per 10µs reading
Sine fit gives RF phase

## Slide 12: Longer term data: 4 minutes of readings at 2Hz
With current setup, correlation with DT is unclear
Vertical X-Y displacement is polarization
Won’t track vibration, acoustics >1Hz
Modulator drive RF has same drift and jitter
Drift: amp heating up?
Jitter: DSP board?
            400fs 1-1M RMS for this board [V. Moore in this workshop]
0
50
100
150
200
250
time, sec
16
14
12
8
6
4
2
0
10
drift 75 fs pp
16
14
12
8
6
4
2
0
10
DN, optical waves
0
50
100
150
200
250
time, sec
X and Y polarization optical phase drift:
Y polarization RF (1GHz) phase drift:
drift 1.6ps
jitter ~300fs pp (~100fs RMS)
1Hz-100kHz
RF phase, milliradians

## Slide 13: Next experiment
laser
DSP
phase mod
ICR
freq. shift
LO
sig.
temperature
controlled
enclosure
New features:
Practical configuration
RX fibers, components in thermally stable box
Loose coil of 100m fiber with heater
Controlled test delay in signal path
Goals:
Demonstrate design parameters
Develop calibration and operation procedures
100m

## Slide 14: Other applications: bunch arrival, length monitor
200ps gaussian pulse input to phase modulator...
I
Q
...exactly reproduced after demodulation:
IF = 500MHz
Commercial LN modulator BW up to 30GHz
Use 60GHz scope or digitizer
rect. to
polar
demodulation

## Slide 15: Other applications: pulsed laser timing
modelocked laser
laser amplifiers
cavity piezo
PLL synchronizer
reference RF (setpoint)
If LLRF loop latency is <10µs, bandwidth is limited by cavity actuator
Reference RF is not independently stabilized
PD
LLRF controls
phase mod
RF monitor receiver
phase mod
master clock

## Slide 16: Conclusions
A time-corrected, CW RF/pulse monitor has been designed, modeled and principles demonstrated
RF amplitude and phase, optical polarization and time delay change have been measured in basic demo
  Ongoing work to reach useful spec
Numerical modeling supports experiments, tests effect of error sources
  Model has shown usefulness for pulse measurements
Theoretically, optical scheme contributes <1fs uncertainty to time delay change measurement
  Receiver hardware needs to be mechanically stabilized
  Digitizer and clock need to be low jitter
Developing concepts for synchronized multichannel operation
Thanks for your attention