## Slide 1
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
1
LLRF Topical Workshop on Advanced Feedback Control
Analysis and Suppression of Ponderomotive
Instabilities in SRF Cavities
Jiayi PENG, Feng QIU
Institute of Modern Physics, Chinese Academy of Sciences
13 October 2026 · Nova Gorica, Slovenia

## Slide 2
Electromechanical Modeling
Summary
Introduction
01
02
04
CONTENTS
Experiments & Suppression
03
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
2

## Slide 3
SRF Cavities
Four cryomodule，23 SRF cavities，cavity bandwidth：80 Hz~200 Hz，HWR010+HWR015）
CAFe：China ADS Front-end (SRF facility, in CW mode)
CAFe (Lanzhou)
CiADS (Huizhou)
  Demonstrate the feasibility of the 10-mA high-power CW proton beam for the China-initiative ADS (CiADS) project
CAFe (China ADS Front-end demo linac) @ IMP
SRF
CiADS
Aerial View
Architectural Rendering
Linac Prototype
HIAF (High Intensity heavy-ion Accl. Facility)
Sep. 2023
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
3

## Slide 4
SRF Cavities
Target：
Highest beam current accelerator for superheavy elements synthesis
Engaging in research on the synthesis of the 119th and 120th element
2022.02, first beam.
2022.03 ~ present, user experiments
Success long-term operation of 8-10 mA CW for 100 hrs
2021.03, stable operation of 10 mA proton beam
2~4 mA
~7 mA
8~10 mA
CAFe: 10 mA proton beam
CAFe2: 0.1 mA Ion beam
China Accelerator Facility for super-heavy new elements (CAFe2)
Spectrometer for Heavy Atoms and
Nuclear Structure
Heavy Ion Source
CAFe beam commissioning
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
4

## Slide 5
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
6
Operational Impact
Growing field and detuning oscillations
Electromechanical coupling can drive self-excited oscillations during CW operation.
The disturbance can involve several cavities and interrupt accelerator operation.
A stability threshold is needed under
actual LLRF feedback settings

## Slide 6
Simplified LFD transfer function
~120 Hz
~120 Hz
0~50 Hz
Ponderomotive Effects
Microphonics
Divergent Oscillation
Vibration (w/o divergency)
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
7
Ponderomotive Effects and Microphonics
Ponderomotive instability: internal electromechanical feedback causes growing oscillations.
Microphonics: external vibration changes the cavity resonance frequency.

## Slide 7
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
8
LLRF Control at CAFE2
Amplitude / phase feedback
Digital PI controllers regulate
the cavity amplitude and phase.
Resonance control
The frequency loop sets the detuning
for the stability measurements.
Both loops remain closed during
the threshold scan.

## Slide 8
Electromechanical Modeling
Summary
Introduction
01
02
04
CONTENTS
Experiments & Suppression
03
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
9

## Slide 9
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
10
Mechanical System Identification
Frequency [Hz]
Grey-box fit to measured detuning
A baseline term plus two resonant modes
reproduces the electromechanical response.
Parameter | Mode 1 | Mode 2
Frequency [Hz] | 118 | 280
Mechanical Q | 15 | 78
Mode strength Gₘ | −3.4 | −1.55
H(s) = Hbase(s) + H₁(s) + H₂(s)

## Slide 10
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
11
Measured and Modeled Detuning
RF turn-off transient
Self-excited oscillation
Oscillation spectrum
The identified model reproduces both transient and oscillatory detuning.

## Slide 11
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
12
Small-signal Model
Cavity electromagnetic dynamics
Perturbations at the operating point
Linearization about the operating point
links field perturbations to detuning.
Normalized variables; negligible beam loading; strongly overcoupled cavity.

## Slide 12
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
13
Closed-loop Stability Criterion
Separate A/P PI controllers
Characteristic equation
Stable: Re(s) < 0 for every characteristic root. Boundary: max Re(s) = 0.

## Slide 13
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
14
Instability Boundaries
Mechanical model variants
K₁: baseline + 118 + 280 Hz
K₂: baseline + 118 Hz
K₃: 280 Hz
Selected operating points
P₁: 280 Hz mode
P₂: 118 Hz mode
P₃: two-mode beating
The excited mechanical mode depends on the operating point.

## Slide 14
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
15
Time-domain Verification
P₁: 280 Hz
P₂: 118 Hz
P₃: two-mode beating
Time traces confirm the mode dependence of the analytical boundaries.

## Slide 15
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
16
Hardware Effects in the Model
Operating-point-dependent loop gain
SSA gain and impedance mismatch enter
the effective cavity and controller parameters.
These corrections connect the ideal model to the measured RF system.

## Slide 16
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
17
Analytical and Simulated Thresholds
Open loop
A/P feedback
Hardware corrections
Characteristic-root analysis and time-domain simulation give consistent onset boundaries.

## Slide 17
Electromechanical Modeling
Summary
Introduction
01
02
04
CONTENTS
Experiments & Suppression
03
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
18

## Slide 18
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
19
Threshold Measurement Procedure
Measurement sequence
1. Set the cavity gradient and PI gains.
2. Sweep toward negative detuning.
3. Slow the scan near the boundary.
4. Fit detuning just before oscillation onset.
Detuning scan near the threshold
Near onset: 0.1–0.15 Hz/s
The threshold is the pre-onset detuning at a fixed cavity gradient.

## Slide 19
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
20
Observed Oscillation Onset
CM4-1
29.8 MV/m
Peak electric field
≈ −37 Hz
Detuning at onset
118 Hz
Dominant mechanical mode

## Slide 20
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
21
Thresholds under Different PI Gains
Gain set 1
Gain set 2
Gain set 3
A: (1, 150)
P: (20, 2000)
A: (1.4, 200)
P: (10, 1000)
A: (0.7, 150)
P: (20, 2000)
A/P entries give (proportional, integral) gains.
Measured thresholds follow the predicted dependence on the PI gains.

## Slide 21
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
22
Verification on a Second Cavity
Two cavities, four gain settings
 | CM4-1 | CM4-2
Half-bandwidth | ≈ 103 Hz | ≈ 113 Hz
Gain sets | 1, 2, 3 | 4
CM4-2: gain set 4
Amplitude PI: (1, 250)
Phase PI: (20, 2000)
The agreement extends to a second cavity and a different RF operating condition.

## Slide 22
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
23
Diagnosis and Controller Retuning
Cause
An attenuation change altered the effective loop gain
without a corresponding controller recalibration.
Diagnosis
The operating point approached the predicted
oscillatory-instability boundary.
Correction
The model guided controller gain adjustment
to recover stable operation.
RF-chain changes require a check of the effective loop gains and stability margin.

## Slide 23
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
24
Suppression in Routine Operation
CM4-1
~17 to < 0.4 faults/day
Gain correction: 23 January 2024
CM4-2
~5 to < 0.4 faults/day
Gain correction: 1 February 2024

## Slide 24
Electromechanical Modeling
Summary
Introduction
01
02
04
CONTENTS
Experiments & Suppression
03
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
25

## Slide 25
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
26
Summary
Mechanical response
The identified modes near 118 and 280 Hz reproduce
measured cavity detuning dynamics.
Stability prediction
The closed-loop characteristic equation connects
PI gains and RF hardware to instability thresholds.
Verification
Two cavities and four gain settings confirm
the predicted dependence of the thresholds.
Suppression
Controller retuning reduced the daily fault rates
from about 17 and 5 to below 0.4.
Stability modeling provides a quantitative basis for reliable LLRF operation.

## Slide 26
Thanks for your attention
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
27
Acknowledgments to the IMP SRF and LLRF teams and collaborators
E-mail: pengjiayi@impcas.ac.cn

## Slide 27
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
28
Backup: Cavity and Mechanical Parameters
HWR015 cavities in CM4
Parameter | Value
RF frequency | 162.5 MHz
Loaded Q | 6 × 10⁵–8 × 10⁵
Half-bandwidth | 100–135 Hz
Peak electric field | ≈ 30 MV/m
R/Q | 382 Ω
V꜀ / Epeak | 0.066 m
Static LFD coefficient | ≈ −0.2 Hz/(MV/m)²
Identified mechanical modes
Parameter | Mode 1 | Mode 2
fₘ [Hz] | 118 | 280
Qₘ | 15 | 78
Gₘ | −3.4 | −1.55
Epeak denotes peak electric field
throughout this presentation.

## Slide 28
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
29
Backup: Characteristic Equations
Open-loop characteristic equation
Closed-loop A/P characteristic equation
The oscillatory boundary corresponds to a root pair at s = ±jΩ.

## Slide 29
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
30
Backup: PI Gain Settings
Set | Cavity | Amplitude (kₚ, kᵢ) | Phase (kₚ, kᵢ)
1 | CM4-1 | (1, 150) | (20, 2000)
2 | CM4-1 | (1.4, 200) | (10, 1000)
3 | CM4-1 | (0.7, 150) | (20, 2000)
4 | CM4-2 | (1, 250) | (20, 2000)
PI controller definitions

## Slide 30
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
31
Backup: A/P and I/Q Control
I/Q characteristic equation
The two controller representations
lead to different coupling terms.
Comparison applies to the stated
model and gain settings.

## Slide 31
J. Peng, Ponderomotive Instabilities in SRF Cavities, TWoAFC 2026, Nova Gorica, 13 October 2026
32
Backup: Impedance Mismatch
Effective cavity parameters
Characteristic equation with hardware corrections
Example for CM4-1: Γ ≈ 0.15 exp(j70°)